Regulatory Considerations in the Design of a Propofol Closed-Loop TCI Pump Controller: Stability

Stephane Bibian, PhD, NeuroWave Systems Inc.,  Guy Dumont, PhD, University of British Columbia

Background/Introduction: The FDA regulatory path for Closed-Loop TCI has been addressed in the 2012 ISAP meeting. Proving the system stability is a critical component of any IDE or PMA application. We propose here a method for the design of a robust predictive PID controller for Propofol whose stability is guaranteed in a given adult population. This work is supported by US Army (W81XWH-11-C-0078).

Methods: Using 44 Propofol PKPD models obtained in a previous study using the WAVcns cortical activity index1,2 on adults (18-60 yrs;53-105 kg), we derived uncertainty bounds using a multiplicative uncertainty framework. These bounds define a limit on the frequency response of the system’s closed-loop transfer function. PID controller parameters can then be optimally tuned so not to violate the stability bounds. A Smith Predictor was further added for improved performance.

Results: We applied the robust predictive controller to all 44 patient models and set the system to target a depth of WAVcns=50. Stable results were obtained for all patients (Figure 1).

Conclusion: A fixed controller that remains stable for a given adult population has been designed. Stability can be mathematically demonstrated. Simulated results show adequate clinical performance.

References

1. Bibian et al., “Patient variability…”, proceedings of the IFAC meeting, 2006

2. Zikov et al., “Quantifying cortical activity…”, IEEE Transactions in Biomedical Engineering, vol.53, n.4, 2006.

0

1

2

3

4

5

6

7

8

9

10